Risk Assessment Form

ID: DR-GGC-RISK-200

Use this form for any detailed risk assessment unless a specific form is provided. Refer to your Summary of Hazards/Risks and complete forms as required, including those that are adequately controlled but could be serious in the absence of active management. The Action Plan and reply section is to help you pursue those requiring action.

| Name of Initial Assessor/Reviewer: | Rosario Lopez | Post Held: | MR Safety Expert |
| --- | --- | --- | --- |
| Department: | Imaging | Date (Initial Review): | 07/03/2023 |
| #### Subject of Assessment: E.g.: hazard, task, equipment, location, people |  |  |  |
| MRI scanning of patients with Implantable Loop recorders (ILR) or insertable cardiac monitors (ICM). Note this risk assessment only applies to 1.5T and 3T and MRI scanning under 6 weeks post-implantation of patients. |  |  |  |
| #### Hazards (Describe the harmful agent(s) and the adverse consequences they could cause) |  |  |  |
| The MRI scanner can present a number of hazards cardiac monitors such as ILR/ICM. What we explicitly review here is any additional risk of MR scanning a patient with an MRI conditional cardiac monitor where the implantation wait is under 6 weeks after implantation. This risk assessment is brought about as some of the manufacturers of MR conditional ILR/ICM typically cite something like ‘At least 6 weeks have elapsed since implantation’ as part of the conditions for safe scanning. The main risk in this case is movement and/or vibration of the device. The 6 week post-implant waiting period allows time for the implant pocket and wound to heal which minimizes the effect of any movement on the device caused by magnetic fields. |  |  |  |
| #### Description of Risk Describe the work that causes exposure to the hazard, and the relevant circumstances. Who is at risk? Highlight significant factors: what makes the risk more or less serious – e.g.: the time taken, how often the work is done, who does it, the work environment, anything else relevant. |  |  |  |
| Apart from the AngelMed Guardian a type of internal monitor that it is MR Unsafe, all the ILR/ICM that we are aware off are MRI Conditional devices. There are four devices that required 6 weeks wait after implantation (Reveal DT 9528 and XT 9529 by Medtronic, and Confirm DM2100 and DM2102 by St Jude Medical). The static magnetic field is not likely to facilitate a significant degree of movement to the MR conditional ILR/ICM as these devices are very light (around 15g or less) and there is little ferromagnetic material in the system and particularly for ILR/ICM the force and torque are minimal. It is difficult to speculate over the potential impact of any potential dislodgement of the device due to the presence of ferrous materials. Moreover, there is no evidence or reported cases of incidents as a result of ILR/ICM being displaced in the MRI environment when patients are scanned within the 6 weeks post implantation period. That said, we do continue to monitor this situation and remain abreast of the latest incidents and developments in this field such that we can adjust our advice rapidly if new evidence was to come to light. |  |  |  |

| Existing Precautions | Describe how they might fail to prevent adverse outcomes. |
| --- | --- |
| On attending their MRI examination, patients are taken through an extensive MRI safety checklist to identify any implants that they may have. MRI safety forums continue to be monitored by MRI Safety experts for any new evidence or incidents that may come to light in this area. The AngelMed Guardian a type of internal cardiac monitor that it is MR Unsafe. | The patient may fail to declare an implant that they have |

Level of Risk - Is the control of this risk adequate?

Give more than one risk level if the assessment covers a range of circumstances. You can use the ‘matrix’ to show how ‘likelihood’ and ‘consequences’ combine to give a conclusion. Also, be critical of existing measures: if you can think how they might fail, or how they could be improved, these are indications of a red or orange risk.

Risk Matrix

| ### Likelihood |  | Impact/Consequences |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | ## Negligible | #### Minor | Moderate | Major | Extreme |
| Almost Certain | Medium | High | High | V High | V High |
| Likely | Medium | Medium | High | High | V High |
| Possible | Low | Medium | Medium | High | High |
| Unlikely | Low | Medium | Medium | Medium | High |
| Rare | Low | Low | Low | Medium | Medium |

Very High                          High                           Medium                      Low

Current risk level

Given the current precautions, and how effective and reliable they are, what is the current level of risk? Green is the target – you have thought it through critically and you have no serious worries. Devise ways of making the risk green wherever you can. Yellow is acceptable but with some reservations. You can achieve these levels by reducing the inherent risk and or by effective and reliable precautions.

High (Orange) or Very High (Red) risks are unacceptable and must be acted on: use the Action Plan section to summarise and communicate the problems and actions required.

Action Plan (if risk level is High (Orange) or Very High (Red)

Use this part of the form for risks that require action.  Use it to communicate, with your Line Manager or Risk Coordinator or others if required.  If using a copy of this form to notify others, they should reply on the form and return to you.  Check that you do receive replies.

Describe the measures required to make the work safe.  Include hardware – engineering controls, and procedures.  Say what you intend to change.  If proposed actions are out with your remit, identify them on the plan below but do not say who or by when; leave this to the manager with the authority to decide this and allocate the resources required.

| Proposed actions to control the problem List the actions required. If action by others is required, you must send them a copy | By Whom | Start date | Action due date |
| --- | --- | --- | --- |

# Action by Others Required - Complete as appropriate: (please tick or enter YES, name and date where appropriate)

| Report up management chain for action |  |
| --- | --- |
| Report to Estates for action |  |
| Contact advisers/specialists |  |
| Alert your staff to problem, new working practice, interim solutions, etc |  |

##### Reply

##### If you receive this form as a manager from someone in your department, you must decide how the risk is to be managed.  Update the action plan and reply with a copy to others who need to know.  If appropriate, you should note additions to the Directorate / Service Risk Register.

If you receive this as an adviser or other specialist, reply to the sender and investigate further as required.

| Date of last review | As per Qpulse record | Next review date | As per Qpulse record |
| --- | --- | --- | --- |